VTRX-FFBD-01 NORMAL_OPERATION
(pump-down)
Power On
Print status
PUMP_DOWN
setup()
Read switch states
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform selfChecks()
- read switch states
- verify valve positioning
- set LED indicators
Raise ERROR + updateLCD()
“R”
Raise ERROR + updateLCD()
“O”
Pressure dose setpoint exceeded
Yes
Failed check?
Sensor OK
Log serial
Transition to: DEBUG_STATE
No
updateLCD()
StartupMsg + updateLCD()
Raise ERROR + updateLCD()
Verify pressure units: mbar
Retry selfChecks()
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- Configure safety relay setpoint
- Verify output type, precision
Verify Pressure safety relay enabled
Verify Pressure Safety Relay CLOSED
Power cut to
High voltage
electronics
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
Print ERROR Expected: CLOSED
Actual: OPEN
No
CLOSED?
Yes
Request
Initial Pressure
Delay
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Yes
Print ERROR
PRNACK
Check if initial pressure approx.
~1013 mbar
Print WARNING
Expected: 1ATM	Actual: <pressure>
No
Within
Threshold?
yes
Log pressure + updateLCD()
Functional Tests
“1.01E3”
*Begin pump-down monitoring*
Delay 1ms
Read switch states
VTRX-BTEST-040
Request pressure + update LCD
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Print Error
 - Argon gate
   valve error
or
 - turbo gate
   valve error
Yes
Contention of valve
positions?
(reading open and
closed simultaneously)
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
Print warning
Open Turbo gate valve!
No
Turbo Gate valve
OPEN?
(closed)
No
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
TODO:
Figure out what to do here when venting to atm
Timeout?
Yes
Yes
No
Log serial + updateLCD()
TODO:
Figure out how to silence this when venting to atm
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Print warning
Close vent valve switch!
Yes
Vent valve
switch OPEN?
Yes
Disable setpoint relay
No
Query Pressure Relay 1 (SS1?) status
Serial log +
Update LCD
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Yes
Command Relay OFF
Log serial + updateLCD()
Vacuum loss error handling threshold
Yes
Print WARNING
Pressure unsafe for high voltage
Pressure above
1E-4 ?
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
No
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify relay still CLOSED
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
Yes
no
CLOSED?
Transition to: DEBUG_STATE
No
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Print ERROR Expected: CLOSED
Actual: OPEN
Yes
Yes
yes
Yes
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
No
No
Transition to: DEBUG_STATE
“M”
Cold cathode failure
PUMPS_POWER_ON?
MicroPirani failure
“C”
TURBO_ROTOR_ON?
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Yes
Within
Threshold?
~1013 mbar
Yes
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
No
Log error (serial + LCD)
Print Warning
Expected: Turbo OFF
Actual: Turbo ON
As expected?
(de-energized)
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
No
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
No
Yes
Pressure below
1E-6?
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Power fed to
High voltage
electronics
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
Yes
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Verify safety relay has OPENED
Yes
ACK?
No
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
Retry limit
exceeded?
yes
Yes
updateLCD()
Print ERROR Expected: OPEN
Actual: CLOSED
Yes
No
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
OPENED ?
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Yes
Transition to: IDLE_STATE
Operational Pressure Reached
Print HVAC STATE
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
yes
No
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Verify safety relay has OPENED
Query relay state to verify
Log SUCCESS +
updateLCD()
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay
VTRX-FFBD-02 STANDARD_VENTING (to Ambient)