VTRX-FFBD-01 NORMAL_OPERATION
(pump-down)
Power On
Log PUMP_DOWN + update LCD
setup()
Read switch states and set
LED indicators
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform selfChecks()
- read power states
- verify valve positioning
- set LED indicators
Raise ERROR + updateLCD()
“R”
Raise ERROR + updateLCD()
“O”
Transition to: DEBUG_STATE
Pressure dose setpoint exceeded
Yes
Sensor OK
Failed check?
Log serial
Transition to: DEBUG_STATE
No
updateLCD()
StartupMsg + updateLCD()
Set pressure units: mbar
Retry selfChecks()
Verify with BW:
- what happens if Turbo rotor ON with PUMPS_POWER OFF?
- Trip safety relay if no successful initial pressure reading?
Raise ERROR + updateLCD()
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- Configure safety relay setpoint
- Verify output type, precision
Verify Pressure safety relay enabled
Verify Pressure safety relay OPEN
Power cut to
High voltage
electronics
Configure pressure sensor safety relay
Request
Initial Pressure
Delay
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Yes
Print ERROR update LCD
Check if initial pressure approx.
~1013 mbar
Trip safety relay
Print WARNING
Expected: 1ATM	Actual: <pressure>
No
Within
Threshold?
yes
Log pressure + updateLCD()
“1.01E3”
*Begin pump-down monitoring*
Delay 1ms
Read switch states and update
LED indicators
Request pressure + update LCD
Print Error
 - Argon gate
   valve error
or
 - turbo gate
   valve error
Yes
Contention of valve
positions?
(i.e. reading open and
closed simultaneously)
No
Print warning
Open Turbo gate valve!
No
Turbo Gate valve
OPEN?
(closed)
Functional Tests
Yes
Print warning
Close vent valve switch!
Yes
Vent valve
switch OPEN?
Perform Pump-down self check
- verify valve positions aren’t contradicting
- read power states
No
Delay
Serial log +
Update LCD
Yes
Yes
Yes
No
LCD Prompt: Close vent valve switch
Failed check?
Vent valve
switch open?
Yes
Turbo Gate valve
OPEN?
PUMPS_POWER_ON?
No
VTRX-BTEST-040
No
No
Yes
updateLCD()
No
No
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
LCD Prompt to
Open gate valve switches
Yes
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Yes
LCD Prompt:
 Turn off pump switches
Log serial
No
TURBO_ROTOR_
ON?
Print Warning?
updateLCD()
No
Yes
retry selfChecks()
Retry limit
exceeded?
No
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Pressure below
1E-6 ?
Pressure above
1E-4 ?
Yes
Raise ERROR + updateLCD()
Vacuum loss error handling threshold
Yes
Print WARNING
Pressure too high
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Timeout?
Timeout?
Yes
Verify relay status
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Command Relay OFF
Log serial + updateLCD()
updateLabview()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Enable Pressure Safety Relay
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
Prompt LCD: turn on PUMPS POWER ON Switch
no
Transition to: DEBUG_STATE
readSwitches()
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
No
Yes
Delay
PUMPS_POWER_ON?
Yes
yes
Yes
No
Yes
No
Switch
Change?
PUMPS
POWER ON?
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
updateLCD()
updateLabview()
Yes
Yes
Transition to: DEBUG_STATE
“M”
Cold cathode failure
MicroPirani failure
“C”
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Within
Threshold?
~1013 mbar
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
Prompt LCD:
Turn Turbo Rotor On
yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
TURBO ROTOR ON?
No
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Yes
updateLCD()
updateLabview()
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Request Pressure
Delay
Yes
ACK?
No
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
Retry limit
exceeded?
yes
Yes
updateLCD()
ACK?
No
No
Retry limit
exceeded?
Yes
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Set SH1 hysteresis: 10%
Yes
Yes
Query setpoint relay direction
Yes
Transition to: IDLE_STATE
Query setpoint hysteresis value
Raise ERROR + updateLCD()
Transition to: IDLE_STATE
updateLabview()
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Print Pressure + updateLabview()
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
yes
No
readSwitches()
BELOW
updateLCD()
Print pressure
No
Enable relay
Verify safety delay enabled**
Start VTRX-FFBD-02
Delay
Set SH1 hysteresis
PUMPS_POWER_ON?
No
Yes
No
Switch
Change?
PUMPS
POWER ON?
Yes
Query relay state to verify
Log SUCCESS +
updateLCD()
No
Yes
TURBO_ROTOR_ON?
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Yes
Log serial + updateLCD
updateLCD()
updateLabview()
Transition to IDLE_STATE
TODO: VTRX-BTEST-900
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay
Delay