Power On
setup()
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Input Pin initialization
Start Logging System
- init errors
Start Logging System
- init errors
Perform selfChecks()
- read switch states
- verify valve positioning
Yes
Failed check?
Log serial
No
updateLCD()
LCD StartupMsg
Retry selfChecks()
Pressure Sensor Configuration
- Config safety relay setpoint
- Enable safety relay
- Verify output type, precision
Set units to mbar
Delay
ACK?
Retry limit
exceeded?
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Set user tag
Request Initial Pressure reading
Check if initial pressure approx.
~1013 mbar
Raise UNEXPECTED_PRESSURE_ERROR
Within
Threshold?
yes
Verify setpoint relay direction: BELOW
Raise PRESSURE_NACK_ERROR
Set Safety relay setpoint: 1.00E-6 mbar
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Query Pressure safety relay status
Raise SAFETY_RELAY_ERROR
ENABLED?
Print status

---
NORMAL_OPERATION
(standard/argon pump-down)
Read switch states
Verify 972b Status
- Transducer status check (@xxxT?;FF)
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Raise ERROR + updateLCD()
“R”
Raise MICROPIRANI_FAILURE
“O”
Pressure dose setpoint exceeded
Sensor OK
SP1 relay error should be persistent and not clearable.
Loop back up to top read Switch States
Raise PRESSURE_DOSE_WARNING
Clear any 972b status errors from LCD queue
Add checks to parse for other errors besides NACK. i.e. incomplete msgs.
Request
Initial Pressure
Raise ERROR
Pressure NACK
If previously had NACK error, and now ACKs, clear previous error
Clear any active pressure NACK error from LCD error queue
Display Pressure
Verify Pressure Safety Relay OPEN?
Pressure below
1E-6 mbar +
threshold?
(HI_VAC)
Power cut to
High voltage
electronics
(PUMP_DOWN)
Raise PRESSURE_RELAY_ERROR Expected: OPEN
Actual: CLOSED
OPEN?
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Verify Pressure Safety Relay CLOSED
Raise SAFETY_RELAY_ERROR Expected: CLOSED
Actual: OPEN
?
CLOSED?
Asynchronous LCD refreshes every 0.5 second with latest pressure and error info
Raise VALVE_CONTENTION_ERROR
Contention of valve
positions?
(reading open and
closed simultaneously)
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Clear any active VALVE_CONTENTION_ERROR from LCD queue
Raise VALVE_CONTENTION_ERROR
Contention of valve
positions?
(reading open and
closed simultaneously)
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
Clear any VALVE_CONTENTION_ERROR from LCD queue
Clear any TURBO_ROTOR_ON_WARNING
Clear any TURBO_ROTOR_ON_WARNING
Raise
ARGON_GATE_VALVE_ERROR
Argon Gate valve
OPEN?
(closed)
Clear any ARGON_GATE_VALVE_ERROR from LCD queue
Raise
VENT_VALVE_OPEN_ERROR
Expected: Vent valve CLOSED
Actual: Vent valve OPEN
Vent Valve Switch
OPEN?
Raise
TURBO_GATE_VALVE_WARNING
Turbo Gate valve
OPEN?
Argon Gate Valve
OPEN?
(closed)
High level things
- This system is wholly dependent on operator actions.
- No automated hardware interlocks prevent operator from progressing to next step if the current step isn’t completed properly.
- Code has to be flexible to this.
- Code has to be very careful in terms of “recommending” actions on LCD. (Serial log can make lots of assumptions though)
Items agnostic to what operator is doing:
- display pressure
- switch malfunction (contention)
- HV relay state verification
- TURBO_ROTOR_ON while PUMPS_POWER OFF
- UNSAFE_FOR_HV messages
Things we should think about removing:
- Argon Gate valve checks
- Turbo gate valve checks
- Vent valve checks
→ these checks entirely depend on the operator is trying to do (unknown to automated system)
- Sequence of venting actions to ATM
1. turbo rotor off
2. PUMPS_POWER_OFF
3. Vent valve switch open (wait for turbomolecular pump to decelerate)
4. Gate valves closed (argon and turbo)
5. Once completed, verify pressure within 1 atm threshold and print message
Clear any TURBO_GATE_VALVE_WARNING from LCD queue
Turbo Gate Valve
OPEN?
Raise
VENT_VALVE_OPEN_ERROR
Vent valve
switch OPEN?
Clear any VENT_VALVE_OPEN_ERROR from LCD queue
Serial log
ValveConfig check OK
Vacuum loss error handling threshold
Raise status
UNSAFE_FOR_HV
Pressure above
1E-4 ?
“unsafe for HV” warning cycles with error messaging
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
Clear any
TURBO_ROTOR_ON_WARNING
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
Clear any
TURBO_ROTOR_ON_WARNING
Ask BW about this behavior?
Remove this
(only check once, conditional on the current pressure)
TODO:
Swap vertical position of turbo rotor check with “pressure above 1E-4”

---
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Set user tag
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Log serial + updateLCD()
Functional Tests
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Log error (serial + LCD)
As expected?
(de-energized)
Command Relay OFF
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify 972b comms
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
no
Transition to: DEBUG_STATE
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Timeout?
Timeout?
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Transition to: DEBUG_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Log error + updateLCD()
Log error + updateLCD()
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
no
Transition to: IDLE_STATE
IS THIS 10%
Setpoint
below 1 atm?
Log + updateLCD()
Verify reading within threshold for 1 atm
Log error + updateLCD()
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
~1 atm
ABOVE
Direction?
Greater than
 +/-10% of setpoint?
BELOW
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay