VTRX-FFBD-01 NORMAL_OPERATION
(standard/argon pump-down)
Power On
Print INIT
Errcount = 0
setup()
DEBUG STATE
Read switch states
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform selfChecks()
- read switch states
- verify valve positioning
Raise ERROR + updateLCD()
“R”
Raise MICROPIRANI_FAILURE
“O”
Pressure dose setpoint exceeded
Yes
Failed check?
Sensor OK
Log serial
No
SP1 relay error should be persistent and not clearable.
Loop back up to top read Switch States
updateLCD()
StartupMsg + updateLCD()
Raise PRESSURE_DOSE_WARNING
Clear any 972b status errors from LCD queue
Retry selfChecks()
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- Config safety relay setpoint
- Enable safety relay
- Verify output type, precision
Query Pressure safety relay enabled
Add checks to parse for other errors besides NACK. i.e. incomplete msgs.
Request
Initial Pressure
Delay
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
Raise ERROR
Relay disabled
No
ENABLED?
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Yes
Yes
Raise ERROR
Pressure NACK
Clear any pressure NACK error from
LCD queue
Check if initial pressure approx.
~1013 mbar
TODO: update error count based on response. Clear prior errors before moving forward
Log WARNING
Expected: 1ATM	Actual: <pressure>
No
Within
Threshold?
Verify Pressure Safety Relay OPEN
No
Pressure above
1E-6 mbar?
yes
(Hi Vac)
Log pressure + updateLCD
Power cut to
High voltage
electronics
Yes
Raise PRESSURE_RELAY_ERROR Expected: OPEN
Actual: CLOSED
No
OPEN?
Yes
If previously had NACK error, and now ACKs, clear previous error
Functional Tests
Clear any PRESSURE_RELAY_ERROR from LCD queue
Verify Pressure Safety Relay CLOSED
TODO: swap relay check with initial pressure request so that the user can see the pressure onscreen
Print ERROR Expected: CLOSED
Actual: OPEN
No
CLOSED?
Asynchronous LCD refreshes every 0.5 second with latest pressure and error info
Yes
Begin pump-down guidance
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Yes
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Request pressure + update LCD
Print Error
 - Argon gate
   valve error
or
 - turbo gate
   valve error
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Contention of valve
positions?
Yes
(reading open and
closed simultaneously)
TODO:
Add check for argon gate valve OPEN
Command Relay OFF
No
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Print warning, Err++
Prompt to open Argon gate
No
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Argon Gate valve
OPEN?
(closed)
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Yes
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
no
Transition to: DEBUG_STATE
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
Yes
yes
Yes
Print warning, Err++
Prompt to open Turbo gate
No
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Turbo Gate valve
OPEN?
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
(closed)
Yes
Yes
Transition to: DEBUG_STATE
“M”
Cold cathode failure
MicroPirani failure
“C”
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Within
Threshold?
~1013 mbar
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
Print warning
Close vent valve switch!
Yes
No
Vent valve
switch OPEN?
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
No
Yes
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
TODO:
Figure out how to silence this when venting to atm
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
Serial log
ValveConfig OK
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Yes
ACK?
No
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
Vacuum loss error handling threshold
Yes
Retry limit
exceeded?
Print Warning
Pressure unsafe for high voltage
yes
Yes
updateLCD()
Pressure above
1E-4 ?
Yes
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Set SH1 hysteresis: 10%
TODO:
Remove relay verification from single loop
No
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
TODO:
Swap vertical position of turbo rotor check with “pressure above 1E-4”
yes
No
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
updateLCD()
No
No
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
Report triggered
Yes
Yes
Log serial + updateLCD
Transition to IDLE_STATE
Print Warning
Expected: Turbo OFF
Actual: Turbo ON
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay
No
Pressure below
1E-6?
Power fed to
High voltage
electronics
Yes
Verify safety relay has OPENED
Print ERROR Expected: OPEN
Actual: CLOSED
No
OPENED ?
Yes
Operational Pressure Reached
Print HVAC STATE
Verify safety relay has OPENED
Remove this
(only check once, conditional on the current pressure)
VTRX-FFBD-02 STANDARD_VENTING (to Ambient)