Power On
setup()
NORMAL_OPERATION
(pump-down)
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Log State change + updateLCD()
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
Perform Pump-down self check
- set LED indicators
- verify valve positions aren’t contradicting
- read power states
Delay
Perform selfChecks()
- read power states
- verify valve positioning
- set LED indicators
Yes
Yes
Yes
No
LCD Prompt: Close vent valve switch
Failed check?
Vent valve
switch open?
Turbo or Roughing
pump ON?
Gate valve switches
OPEN?
Yes
Failed check?
No
Log serial
No
No
Yes
No
updateLCD()
updateLCD()
StartupMsg + updateLCD()
LCD Prompt to
Open gate valve switches
Retry selfChecks()
LCD Prompt:
 Turn off pump switches
Log serial
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- output type, precision,etc.
Verify 972b Status
- Transducer status check (@xxxT?;FF)
updateLCD()
No
retry selfChecks()
Retry limit
exceeded?
Configure pressure sensor safety relay
“M”
Cold cathode failure
Yes
MicroPirani failure
“C”
status?
Raise ERROR + updateLCD()
Raise ERROR + updateLCD()
“R”
Raise ERROR + updateLCD()
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
“O”
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Transition to: IDLE_STATE
Set pressure units: mbar
Request Pressure
Delay
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Yes
TBD: IDLE_STATE
Raise ERROR + updateLCD()
updateLabview()
Raise ERROR + updateLCD()
Log + updateLCD()
Print Pressure
“1.01E3”
Check Valve Pos.
- Gate valve switches: OPEN
- Vent valve switch: CLOSED
Yes
Verify initial pressure
~1.0E00 mbar
Retry limit
exceeded?
No
Perform selfChecks()
- read power states
- verify valve positions
- set LED indicators
Raise ERROR
Request calibration
No
Within
Threshold?
yes
Yes
Failed check?
Log serial
Log pressure + updateLCD()
972b sensor DEFAULT settings
Safety relay value = 1.00E0 Torr = 1333 mbar
Safety relay hysteresis = 1.10E0 Torr =
Safety relay direction = BELOW
No
updateLCD()
updateLabview()
Log serial + updateLCD()
Transition to REMOTE_CONTROL
State
Enable Pressure Safety Relay
Request for
REMOTE ctl?
No
Prompt LCD: turn on Roughing Pump Switch
Read Pressure
readSwitches()
No
Log pressure + updateLCD()
Delay
PUMPS_POWER_ON?
Switch
Change?
Switch
Change?
Delay
No
Yes
updateLCD()
updateLabview()
Prompt LCD:
Turn Turbo Rotor On
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Functional Tests
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Timeout?
Timeout?
Yes
TODO: VTRX-BTEST-050
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Yes
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
No
Log error (serial + LCD)
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
As expected?
(de-energized)
Yes
Command Relay OFF
Log serial + updateLCD()
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Transition to: IDLE_STATE
Timeout?
Timeout?
Yes
VTRX-BTEST-020
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Test serial reading from pressure gauge using arduino at 1 atm
Log error + updateLCD()
Yes
Request Pressure
no
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Transition to: IDLE_STATE
ACK?
no
Retry limit
exceeded?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
No
Log error (serial + LCD)
yes
Yes
As expected?
(de-energized)
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Yes
Retry limit
exceeded?
Command Relay OFF
Log serial + updateLCD()
Transition to: IDLE_STATE
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform calibration
Transition to: IDLE_STATE
Within
Threshold?
~1013 mbar
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Transition to: IDLE_STATE
Query setpoint switch value (SP1?)
Query setpoint switch value (SP1?)
No
No
No
No
Setpoint
Below 1 atm?
Set pressure units: mbar
Setpoint
below 1 atm?
Log + updateLCD()
Loop
Timeout?
Log + updateLCD()
Retry limit
exceeded?
Yes
Yes
Yes
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Watchdog timer
Query setpoint hysteresis value
Watchdog timer
Log error + updateLCD()
Transition to: IDLE_STATE
Query setpoint hysteresis value
Request Pressure
no
Transition to: IDLE_STATE
Yes
ACK?
Yes
Greater than
 +/-10% of setpoint?
no
Greater than
 +/-10% of setpoint?
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
Log serial
No
Yes
No
yes
Yes
updateLCD()
updateLCD()
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Configure setpoint safety delay**
Retry limit
exceeded?
Set SH1 hysteresis: 10%
Set SH1 hysteresis
Query setpoint relay direction
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: ABOVE
Within
Threshold?
~1 atm
BELOW
Direction?
Enable 972b setpoint
yes
ABOVE
Log serial + updateLCD
Print pressure
Enable relay
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Verify Relay Enabled
Query relay state to verify
Log SUCCESS +
updateLCD()
Log serial + UpdateLCD()
Transition to: IDLE_STATE
TODO: above/below
Transition to IDLE_STATE
TODO: VTRX-BTEST-900