Power On
setup()
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Input Pin initialization
Start Logging System
- init errors
Start Logging System
- init errors
Read switch states
Pressure Sensor Configuration
- Config safety relay setpoint
- Enable safety relay
- Verify output type, precision
Set units to mbar
Delay
ACK?
No
Retry limit
exceeded?
Yes
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Set user tag
Request Initial Pressure reading
Check if initial pressure approx.
~1013 mbar
Raise UNEXPECTED_PRESSURE_ERROR
Within
Threshold?
Consider an
Auto-reset for this error
after some amount
of time: 10 minutes
yes
Verify setpoint relay direction: BELOW
Loop NACK retries back up to re-read
Pressure, not just to previous command
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Set Safety relay setpoint: 1.00E-4 mbar
BW says change this to
1.00E-4 mbar
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Verify Pressure safety relay enabled
    ACK?
Raise PRESSURE_NACK_ERROR
Print status
SETUP COMPLETE
Verify
errorcount == 0
Transition to
NORMAL
Add turbo rotor check her as well as main loop

---
NORMAL_OPERATION
(standard/argon pump-down)
Read system status
Verify 972b Status
- Transducer status check (@xxxT?;FF)
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Raise ERROR + updateLCD()
“R”
Raise MICROPIRANI_FAILURE
“O”
Pressure dose setpoint exceeded
Sensor OK
SP1 relay error should be persistent and not clearable.
Loop back up to top read Switch States
Raise PRESSURE_DOSE_WARNING
Clear MicroPirani or ColdCathode status errors from LCD queue
Add checks to parse for other errors besides NACK. i.e. incomplete msgs.
Request
Initial Pressure
RETRY LIMIT:3
Raise ERROR
Pressure NACK
Clear any active PRESSURE_NACK_ERROR from LCD error queue
If previously sensor NACK’s, and now ACKs, clear previous error
Display Pressure
Update pressure in this screenshot
Verify Pressure Safety Relay OPEN?
Pressure below
1E-4 mbar +
threshold?
(HI_VAC)
If RELAY_ERROR has been
exists for >1 sec,
Make the error persistent
Power cut to
High voltage
electronics
Asynchronous LCD refreshes every 0.5 second with latest pressure and error info
(PUMP_DOWN)
Raise PRESSURE_RELAY_ERROR Expected: OPEN
Actual: CLOSED
OPEN?
Raise status
UNSAFE_FOR_HV
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Verify Pressure Safety Relay CLOSED
Clear any UNSAFE_FOR_HV status
- Sequence of venting actions to ATM
1. turbo rotor off
2. PUMPS_POWER_OFF
3. Vent valve switch open (wait for turbomolecular pump to decelerate)
4. Gate valves closed (argon and turbo)
5. Once completed, verify pressure within 1 atm threshold and print message
Raise SAFETY_RELAY_ERROR Expected: CLOSED
Actual: OPEN
Raise VALVE_CONTENTION_ERROR
CLOSED?
Contention of valve
positions?
(reading open and
closed simultaneously)
If this error has been existing for >1 sec,
Make the error persistent
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Clear any active VALVE_CONTENTION_ERROR from LCD queue
“unsafe for HV” warning cycles with error messaging
Raise VALVE_CONTENTION_ERROR
Contention of valve
positions?
(reading open and
closed simultaneously)
Clear any VALVE_CONTENTION_ERROR from LCD queue
Vacuum loss error handling threshold
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
TURBO_ROTOR_ON?
PUMPS_POWER_ON?
High level things
- This system is wholly dependent on operator actions.
- No automated hardware interlocks prevent operator from progressing to next step if the current step isn’t completed properly.
- Code has to be flexible to this.
- Code has to be very careful in terms of “recommending” actions on LCD. (Serial log can make lots of assumptions though)
Items agnostic to what operator is doing:
- display pressure
- switch malfunction (contention)
- HV relay state verification
- TURBO_ROTOR_ON while PUMPS_POWER OFF
- UNSAFE_FOR_HV messages
Things we should think about removing:
- Argon Gate valve checks
- Turbo gate valve checks
- Vent valve checks
→ these checks entirely depend on the operator is trying to do (unknown to automated system)
Clear any TURBO_ROTOR_ON_WARNING
Clear any TURBO_ROTOR_ON_WARNING
Clear any
TURBO_ROTOR_ON_WARNING
Clear any
TURBO_ROTOR_ON_WARNING
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
Delay 1 ms
Delay 1 ms

---
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Set user tag
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Log serial + updateLCD()
Functional Tests
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Log error (serial + LCD)
As expected?
(de-energized)
Command Relay OFF
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify 972b comms
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
no
Transition to: DEBUG_STATE
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Timeout?
Timeout?
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Transition to: DEBUG_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Log error + updateLCD()
Log error + updateLCD()
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
no
Log serial
Transition to: IDLE_STATE
IS THIS 10%
Setpoint
below 1 atm?
Log + updateLCD()
updateLCD()
Verify reading within threshold for 1 atm
Log error + updateLCD()
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
~1 atm
ABOVE
Direction?
Greater than
 +/-10% of setpoint?
Log serial
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
updateLCD()
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay