VTRX-FFBD-01 NORMAL_OPERATION
(standard/argon pump-down)
Power On
Set units to mbar
setup()
Delay
Read switch states
ACK?
No
No
Retry limit
exceeded?
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Yes
Yes
Yes
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Raise ERROR
Pressure NACK
Input Pin config
Clear any PRESSURE_NACK_ERROR
from LCD queue
Start Logging System
- init errors
Start Logging System
- init errors
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Set user tag
Perform selfChecks()
- read switch states
- verify valve positioning
Delay
Raise ERROR + updateLCD()
“R”
Raise MICROPIRANI_FAILURE
“O”
ACK?
No
No
Retry limit
exceeded?
Pressure dose setpoint exceeded
Yes
Failed check?
Sensor OK
Log serial
Yes
Yes
Yes
No
SP1 relay error should be persistent and not clearable.
Loop back up to top read Switch States
Raise ERROR
Pressure NACK
updateLCD()
Clear any PRESSURE_NACK_ERROR
from LCD queue
StartupMsg + updateLCD()
Raise PRESSURE_DOSE_WARNING
Clear any 972b status errors from LCD queue
Retry selfChecks()
Pressure Sensor Configuration
- Config safety relay setpoint
- Enable safety relay
- Verify output type, precision
Add checks to parse for other errors besides NACK. i.e. incomplete msgs.
Request
Initial Pressure
Delay
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Yes
Raise ERROR
Pressure NACK
Clear any pressure NACK error from
LCD queue
Check if initial pressure approx.
~1013 mbar
Display Pressure
TODO: update error count based on response. Clear prior errors before moving forward
Log WARNING
Expected: 1ATM	Actual: <pressure>
No
Within
Threshold?
yes
Verify Pressure Safety Relay OPEN?
Yes
Log pressure + updateLCD
Pressure below
1E-6 mbar?
(HI_VAC)
Power cut to
High voltage
electronics
No
(PUMP_DOWN)
Raise PRESSURE_RELAY_ERROR Expected: OPEN
Actual: CLOSED
No
Query Pressure safety relay enabled
OPEN?
If previously had NACK error, and now ACKs, clear previous error
Yes
Clear any PRESSURE_RELAY_ERROR from LCD queue
Raise ERROR
Relay disabled
No
Verify Pressure Safety Relay CLOSED
ENABLED?
TODO: swap relay check with initial pressure request so that the user can see the pressure onscreen
Yes
Raise status
SAFE_FOR_HV
Raise SAFETY_RELAY_ERROR Expected: CLOSED
Actual: OPEN
Yes
No
CLOSED?
Asynchronous LCD refreshes every 0.5 second with latest pressure and error info
Yes
TURBO_ROTOR_ON?
Clear any active PRESSURE_RELAY_ERROR from LCD queue
No
VTRX-BTEST-040
Raise VALVE_CONTENTION_ERROR
Yes
Contention of valve
positions?
PUMPS_POWER_ON?
(reading open and
closed simultaneously)
No
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Clear any VALVE_CONTENTION_ERROR from LCD queue
Raise status
SAFE_FOR_HV
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Raise
ARGON_GATE_VALVE_ERROR
No
Argon Gate valve
OPEN?
(closed)
Yes
No
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Clear any ARGON_GATE_VALVE_ERROR from LCD queue
Yes
No
Log serial + updateLCD()
Functional Tests
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Disable setpoint relay
Raise
TURBO_GATE_VALVE_WARNING
Query Pressure Relay 1 (SS1?) status
No
Turbo Gate valve
OPEN?
(closed)
Yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Clear any TURBO_GATE_VALVE_WARNING from LCD queue
Command Relay OFF
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Raise
VENT_VALVE_OPEN_ERROR
Yes
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Vent valve
switch OPEN?
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
TODO:
Figure out how to silence this when venting to atm
Request Pressure
No
no
Transition to: DEBUG_STATE
Clear any VENT_VALVE_OPEN_ERROR from LCD queue
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
Yes
yes
Yes
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Serial log
ValveConfig check OK
Yes
Transition to: DEBUG_STATE
“M”
Cold cathode failure
MicroPirani failure
“C”
Vacuum loss error handling threshold
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Within
Threshold?
~1013 mbar
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Raise status
UNSAFE_FOR_HV
Yes
TODO:
Remove relay verification from single loop
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Pressure above
1E-4 ?
“unsafe for HV” warning cycles with error messaging
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
No
Log serial + updateLCD()
No
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
No
No
Yes
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Yes
Request Pressure
Yes
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Clear any
TURBO_ROTOR_ON_WARNING
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
Yes
ACK?
No
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
Retry limit
exceeded?
yes
Yes
updateLCD()
Yes
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Delay
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
yes
No
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
updateLCD()
TODO:
Swap vertical position of turbo rotor check with “pressure above 1E-4”
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay
Remove this
(only check once, conditional on the current pressure)
VTRX-FFBD-02 STANDARD_VENTING (to Ambient)