VTRX-FFBD-01 NORMAL_OPERATION
(pump-down)
Power On
Log PUMP_DOWN + update LCD
setup()
Read switch states and set
LED indicators
Verify with BW:
- what happens if Turbo rotor ON with PUMPS_POWER OFF?
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform selfChecks()
- read power states
- verify valve positioning
- set LED indicators
Raise ERROR + updateLCD()
“R”
Raise ERROR + updateLCD()
“O”
Transition to: DEBUG_STATE
Pressure dose setpoint exceeded
Yes
Failed check?
Sensor OK
Log serial
Transition to: DEBUG_STATE
No
updateLCD()
StartupMsg + updateLCD()
Set pressure units: mbar
Retry selfChecks()
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- Configure safety relay setpoint
- Verify output type, precision
Verify Pressure safety relay enabled
Verify Pressure Safety Relay OPEN
Power cut to
High voltage
electronics
Configure pressure sensor safety relay
OPEN?
No
Print ERROR update LCD
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
Yes
Request
Initial Pressure
Delay
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Yes
Print ERROR update LCD
Check if initial pressure approx.
~1013 mbar
Trip safety relay
Print WARNING
Expected: 1ATM	Actual: <pressure>
No
Within
Threshold?
yes
Log pressure + updateLCD()
Functional Tests
“1.01E3”
*Begin pump-down monitoring*
Delay 1ms
Read switch states and update
LED indicators
VTRX-BTEST-040
Request pressure + update LCD
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Print Error
 - Argon gate
   valve error
or
 - turbo gate
   valve error
Yes
Contention of valve
positions?
(i.e. reading open and
closed simultaneously)
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
Print warning
Open Turbo gate valve!
No
Turbo Gate valve
OPEN?
(closed)
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Timeout?
Timeout?
Yes
Yes
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Print warning
Close vent valve switch!
Yes
Vent valve
switch OPEN?
Yes
Disable setpoint relay
No
Query Pressure Relay 1 (SS1?) status
Serial log +
Update LCD
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Yes
Command Relay OFF
Log serial + updateLCD()
No
No
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
Yes
VTRX-BTEST-020
Yes
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
No
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
TURBO_ROTOR_
ON?
Print Warning?
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
Yes
no
No
No
Transition to: DEBUG_STATE
Pressure below
1E-6 ?
Pressure above
1E-4 ?
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
Yes
Vacuum loss error handling threshold
Yes
Yes
yes
Yes
Print WARNING
Pressure unsafe for high voltage
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Transition to: DEBUG_STATE
“M”
Cold cathode failure
MicroPirani failure
“C”
Verify relay still open
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Within
Threshold?
~1013 mbar
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Operational Pressure reached
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
No
Power fed to
High voltage
electronics
Verify safety relay has CLOSED
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Yes
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Yes
ACK?
No
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
VTRX-FFBD-02 STANDARD_VENTING (to Ambient)
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
Retry limit
exceeded?
yes
Yes
updateLCD()
Yes
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
yes
No
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
updateLCD()
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay