Power On
NORMAL_OPERATION
(pump-down)
setup()
INITIALIZATION STATE
init SystemStates
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
set currentState = INITIALIZATION_STATE
Log State change + updateLCD()
Input Pin config
Perform Pump-down selfChecks()
- set LED indicators
- verify valve positions aren’t contradicting
- read power states
Start Logging System
- init errors
Start Logging System
- init errors
Main Loop
Delay
Perform selfChecks()
- read power states
- verify valve positioning
- set LED indicators
switch (currentState)
Yes
Yes
Yes
No
LCD Prompt: Close vent valve switch
Failed check?
Vent valve
switch open?
Turbo or Roughing
pump ON?
Gate valve switches
open?
No
No
No
Yes
Yes
Failed check?
Log serial
updateLCD()
No
updateLCD()
LCD Prompt:
Open gate valve switches
StartupMsg + updateLCD()
LCD Prompt:
 Turn off pump switches
Log serial
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Retry selfChecks()
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- output type, precision,etc.
updateLCD()
No
retry selfChecks()
Retry limit
exceeded?
“M”
Cold cathode failure
Yes
MicroPirani failure
“C”
status?
Configure pressure sensor safety relay
Raise ERROR + updateLCD()
Raise ERROR + updateLCD()
“R”
Raise ERROR + updateLCD()
“O”
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
IDLE_STATE
Transition to: IDLE_STATE
Log + updateLCD()
Set pressure units: mbar
Check Valve Pos.
- Gate valve switches: OPEN
- Vent valve switch: CLOSED
Request Pressure
Delay
DEFAULTS
Safety relay value = 1.00E0 Torr = 1333 mbar
ACK?
No
No
Retry limit
exceeded?
Perform selfChecks()
- read power states
- verify valve positions
- set LED indicators
Yes
Yes
Yes
Raise ERROR + updateLCD()
updateLabview()
Raise ERROR + updateLCD()
Yes
Failed check?
Log serial
Print Pressure
No
updateLCD()
“1.01E3”
Log serial + updateLCD()
Yes
Verify initial pressure
~1.0E00 mbar
No
Retry limit
exceeded?
Transition to REMOTE_CONTROL
State
Request for
REMOTE ctl?
No
Raise ERROR
Request calibration
no
Within
Threshold?
~1 atm
Read Pressure
yes
Log pressure + updateLCD()
Log pressure + updateLCD()
Delay
updateLabview()
Enable Pressure Safety Relay
Prompt LCD: turn on Roughing Pump Switch
Delay
Pump On?
Timeout?
Functional Tests
updateLCD()
updateLabview()
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Timeout?
Timeout?
Yes
No
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
No
Log error (serial + LCD)
As expected?
(de-energized)
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Yes
Command Relay OFF
Log serial + updateLCD()
Transition to: IDLE_STATE
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Timeout?
Timeout?
Yes
VTRX-BTEST-020
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Log error + updateLCD()
Request Pressure
Test serial reading from pressure gauge using arduino at 1 atm
Yes
no
Disable setpoint relay
Transition to: IDLE_STATE
Query Pressure Relay 1 (SS1?) status
ACK?
no
Retry limit
exceeded?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
yes
Yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Yes
Command Relay OFF
Transition to: IDLE_STATE
Log serial + updateLCD()
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform calibration
Within
Threshold?
~1013 mbar
Transition to: IDLE_STATE
No
Log error + updateLCD()
“R”
Log error + updateLCD()
yes
“O”
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Query setpoint switch value (SP1?)
No
No
No
Setpoint
Below 1 atm?
Log + updateLCD()
Loop
Timeout?
Set pressure units: mbar
No
Setpoint
below 1 atm?
Log + updateLCD()
Retry limit
exceeded?
Yes
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Yes
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Watchdog timer
Query setpoint hysteresis value
Watchdog timer
Log error + updateLCD()
Request Pressure
Transition to: IDLE_STATE
Query setpoint hysteresis value
no
Yes
Transition to: IDLE_STATE
Greater than
 +/-10% of setpoint?
Log serial
Transition to: IDLE_STATE
ACK?
no
Yes
Retry limit
exceeded?
Greater than
 +/-10% of setpoint?
Log serial
No
updateLCD()
Yes
yes
Yes
No
updateLCD()
Verify reading within threshold for 1 atm
Log error + updateLCD()
Set SH1 hysteresis: 10%
No
Query setpoint relay direction
Configure setpoint safety delay**
Retry limit
exceeded?
Set SH1 hysteresis
Transition to: IDLE_STATE
Command setpoint relay direction: ABOVE
BELOW
Perform calibration
no
Within
Threshold?
~1 atm
Direction?
Enable 972b setpoint
ABOVE
yes
Enable relay
Print pressure
Log serial + updateLCD
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Query relay state to verify
Verify Relay Enabled
Log SUCCESS +
updateLCD()
Log serial + UpdateLCD()
Transition to: IDLE_STATE
TODO: VTRX-BTEST-900
Transition to IDLE_STATE
TODO: above/below