Power On
setup()
NORMAL_OPERATION
(standard/argon pump-down)
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Input Pin config
Read switch states
Start Logging System
- init errors
Start Logging System
- init errors
Perform selfChecks()
- read switch states
- verify valve positioning
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
Failed check?
Log serial
“M”
Cold cathode failure
MicroPirani failure
“C”
No
status?
updateLCD()
LCD StartupMsg
Raise ERROR + updateLCD()
“R”
Raise MICROPIRANI_FAILURE
“O”
Retry selfChecks()
Pressure Sensor Configuration
- Config safety relay setpoint
- Enable safety relay
- Verify output type, precision
Pressure dose setpoint exceeded
Sensor OK
SP1 relay error should be persistent and not clearable.
Loop back up to top read Switch States
Raise PRESSURE_DOSE_WARNING
Clear any 972b status errors from LCD queue
Add checks to parse for other errors besides NACK. i.e. incomplete msgs.
Request
Initial Pressure
Delay
ACK?
Retry limit
exceeded?
Set units to mbar
Raise ERROR
Pressure NACK
Clear any active pressure NACK error from LCD error queue
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Display Pressure
TODO: update error count based on response. Clear prior errors before moving forward
Set user tag
Verify Pressure Safety Relay OPEN?
Pressure below
1E-6 mbar +
threshold?
Request Initial Pressure reading
(HI_VAC)
Power cut to
High voltage
electronics
(PUMP_DOWN)
Raise PRESSURE_RELAY_ERROR Expected: OPEN
Actual: CLOSED
Check if initial pressure approx.
~1013 mbar
OPEN?
If previously had NACK error, and now ACKs, clear previous error
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Raise UNEXPECTED_PRESSURE_ERROR
Verify Pressure Safety Relay CLOSED
Within
Threshold?
yes
Raise SAFETY_RELAY_ERROR Expected: CLOSED
Actual: OPEN
Verify setpoint relay direction: BELOW
?
CLOSED?
Asynchronous LCD refreshes every 0.5 second with latest pressure and error info
Raise VALVE_CONTENTION_ERROR
Contention of valve
positions?
(reading open and
closed simultaneously)
Clear any active VALVE_CONTENTION_ERROR from LCD queue
Raise PRESSURE_NACK_ERROR
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
Set Safety relay setpoint: 1.00E-6 mbar
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
Clear any VALVE_CONTENTION_ERROR from LCD queue
Clear any TURBO_ROTOR_ON_WARNING
Vent Valve Switch
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
ARGON_GATE_VALVE_ERROR
Argon Gate valve
(closed)
Clear any ARGON_GATE_VALVE_ERROR from LCD queue
Query Pressure safety relay status
VENT_VALVE_OPEN_ERROR
Raise SAFETY_RELAY_ERROR
ENABLED?
TURBO_GATE_VALVE_WARNING
Turbo Gate valve
Argon Gate Valve
- Sequence of venting actions to ATM
1. turbo rotor off
2. PUMPS_POWER_OFF
3. Vent valve switch open (wait for turbomolecular pump to decelerate)
4. Gate valves closed (argon and turbo)
5. Once, completed, and pressure within limit,
Clear any TURBO_GATE_VALVE_WARNING from LCD queue
Argon Turbo Gate
Vent valve
switch OPEN?
Clear any VENT_VALVE_OPEN_ERROR from LCD queue
Serial log
ValveConfig check OK
Vacuum loss error handling threshold
Raise status
UNSAFE_FOR_HV
Pressure above
1E-4 ?
“unsafe for HV” warning cycles with error messaging
Clear any
Ask BW about this behavior?
Remove this
(only check once, conditional on the current pressure)
TODO:
Swap vertical position of turbo rotor check with “pressure above 1E-4”

---
TODO: update error count based on response. Clear prior errors before moving forward

---
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Set user tag
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Log serial + updateLCD()
Functional Tests
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Log error (serial + LCD)
As expected?
(de-energized)
Command Relay OFF
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify 972b comms
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
no
Transition to: DEBUG_STATE
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Timeout?
Timeout?
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Transition to: DEBUG_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Log error + updateLCD()
Log error + updateLCD()
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
no
Transition to: IDLE_STATE
IS THIS 10%
Setpoint
below 1 atm?
Log + updateLCD()
Verify reading within threshold for 1 atm
Log error + updateLCD()
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
~1 atm
ABOVE
Direction?
Greater than
 +/-10% of setpoint?
BELOW
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay