VTRX-FFBD-01 NORMAL_OPERATION
(pump-down)
Power On
Log State change + updateLCD()
setup()
Read switch input and set LED indicators
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform selfChecks()
- read power states
- verify valve positioning
- set LED indicators
Raise ERROR + updateLCD()
“R”
Raise ERROR + updateLCD()
“O”
Transition to: DEBUG_STATE
Pressure dose setpoint exceeded
Yes
Sensor OK
Failed check?
Log serial
Transition to: DEBUG_STATE
No
updateLCD()
StartupMsg + updateLCD()
Set pressure units: mbar
Retry selfChecks()
Raise ERROR + updateLCD()
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- output type, precision,etc.
Request Pressure
Delay
ACK?
No
Configure pressure sensor safety relay
No
Retry limit
exceeded?
Yes
Yes
Yes
Raise ERROR + updateLCD()
updateLabview()
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
Check if initial pressure approx.
~1013 mbar
972b sensor DEFAULT settings
Safety relay value = 1.00E0 Torr = 1333 mbar
Safety relay hysteresis = 1.10E0 Torr =
Safety relay direction = BELOW
Print warning
Expected: 1ATM	Actual: <pressure>
No
Within
Threshold?
yes
Log pressure + updateLCD()
“1.01E3”
Begin basic self checks
Print Warning
 - Argon gate
   valve error
or
 - turbo gate
   valve error
Yes
Contradicting valve
position?
No
Print warning
Open gate valve switches
No
Turbo Gate valve
OPEN?
Yes
Yes
Print warning
Vent valve switch open!
Vent valve
switch open?
LCD Prompt: Close vent valve switch
Functional Tests
Perform Pump-down self check
- verify valve positions aren’t contradicting
- read power states
Delay
Yes
Yes
Yes
No
LCD Prompt: Close vent valve switch
Failed check?
Vent valve
switch open?
Turbo Gate valve
OPEN?
PUMPS_POWER_ON?
No
No
No
Yes
updateLCD()
LCD Prompt to
Open gate valve switches
VTRX-BTEST-040
LCD Prompt:
 Turn off pump switches
Log serial
updateLCD()
No
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
retry selfChecks()
Retry limit
exceeded?
Yes
Raise ERROR + updateLCD()
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
updateLabview()
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Timeout?
Timeout?
Yes
Enable Pressure Safety Relay
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Prompt LCD: turn on PUMPS POWER ON Switch
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
readSwitches()
No
No
Log error (serial + LCD)
As expected?
(de-energized)
Delay
PUMPS_POWER_ON?
No
Yes
No
Yes
Switch
Change?
PUMPS
POWER ON?
Command Relay OFF
Log serial + updateLCD()
updateLCD()
updateLabview()
Yes
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Prompt LCD:
Turn Turbo Rotor On
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
no
Transition to: DEBUG_STATE
TURBO ROTOR ON?
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
Yes
yes
Yes
updateLCD()
updateLabview()
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Transition to: DEBUG_STATE
“M”
Cold cathode failure
MicroPirani failure
“C”
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Request Pressure
Delay
Within
Threshold?
~1013 mbar
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
No
Log error (serial + LCD)
ACK?
No
As expected?
(de-energized)
No
Retry limit
exceeded?
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Yes
Yes
Yes
Command Relay OFF
Log serial + updateLCD()
Raise ERROR + updateLCD()
updateLabview()
No
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Print Pressure + updateLabview()
Yes
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
readSwitches()
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
No
Start VTRX-FFBD-02
no
Delay
Query setpoint switch value (SP1?)
PUMPS_POWER_ON?
Transition to: IDLE_STATE
No
Yes
No
Switch
Change?
PUMPS
POWER ON?
Yes
ACK?
No
Yes
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
No
Retry limit
exceeded?
Yes
yes
Yes
updateLCD()
TURBO_ROTOR_ON?
Yes
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Yes
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
updateLCD()
updateLabview()
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
yes
No
BELOW
Delay
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
updateLCD()
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
TODO: VTRX-BTEST-900
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay