VTRX-FFBD-01 NORMAL_OPERATION
(pump-down)
Power On
Print status
PUMP_DOWN
setup()
Read switch states
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Perform selfChecks()
- read switch states
- verify valve positioning
- set LED indicators
Raise ERROR + updateLCD()
“R”
Raise ERROR + updateLCD()
“O”
Pressure dose setpoint exceeded
Yes
Failed check?
Sensor OK
Log serial
Transition to: DEBUG_STATE
No
updateLCD()
StartupMsg + updateLCD()
Raise ERROR + updateLCD()
Verify pressure units: mbar
Retry selfChecks()
Pressure Sensor Configuration
- Set units to mbar
- Configure user tag
- Configure safety relay setpoint
- Verify output type, precision
Verify Pressure safety relay enabled
Verify Pressure Safety Relay CLOSED
Power cut to
High voltage
electronics
Transition to
TEST_STATE/NORMAL_OPERATION
Transition to
TEST_STATE/NORMAL_OPERATION
Print ERROR Expected: CLOSED
Actual: OPEN
No
CLOSED?
Yes
Request
Initial Pressure
Delay
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Yes
Print ERROR
PRNACK
Check if initial pressure approx.
~1013 mbar
Print WARNING
Expected: 1ATM	Actual: <pressure>
No
Within
Threshold?
yes
Log pressure + updateLCD()
Functional Tests
*Begin pump-down monitoring*
Delay 1ms
Read switch states
VTRX-BTEST-040
Request pressure + update LCD
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Print Error
 - Argon gate
   valve error
or
 - turbo gate
   valve error
Yes
Contention of valve
positions?
(reading open and
closed simultaneously)
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
TODO:
Figure out how to clear this once operator starts venting to atm
Print warning
Open Turbo gate valve!
No
Turbo Gate valve
OPEN?
(closed)
No
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Yes
Yes
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Print warning
Close vent valve switch!
Yes
Vent valve
switch OPEN?
Yes
Disable setpoint relay
No
Query Pressure Relay 1 (SS1?) status
TODO:
Figure out how to silence this when venting to atm
Serial log +
Update LCD
Lasts for 2 seconds then resets{
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Yes
Command Relay OFF
Log serial + updateLCD()
Vacuum loss error handling threshold
Yes
Print reminder
Pressure unsafe for high voltage
Pressure above
1E-4 ?
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
No
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify relay still CLOSED
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
Yes
no
CLOSED?
Transition to: DEBUG_STATE
No
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Print ERROR Expected: CLOSED
Actual: OPEN
Yes
Yes
yes
Yes
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
No
No
Transition to: DEBUG_STATE
“M”
Cold cathode failure
PUMPS_POWER_ON?
MicroPirani failure
“C”
TURBO_ROTOR_ON?
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Yes
Within
Threshold?
~1013 mbar
Yes
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
No
Log error (serial + LCD)
Print Warning
Expected: Turbo OFF
Actual: Turbo ON
As expected?
(de-energized)
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
No
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
No
Yes
Pressure below
1E-6?
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Power fed to
High voltage
electronics
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
Yes
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Verify safety relay has OPENED
Yes
ACK?
No
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
Retry limit
exceeded?
yes
Yes
updateLCD()
Print ERROR Expected: OPEN
Actual: CLOSED
Yes
No
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
OPENED ?
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Yes
Transition to: IDLE_STATE
Operational Pressure Reached
Print HVAC STATE
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
yes
No
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Verify safety relay has OPENED
Query relay state to verify
Log SUCCESS +
updateLCD()
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay
VTRX-FFBD-02 STANDARD_VENTING (to Ambient)