Power On
setup()
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
NORMAL_OPERATION
(standard/argon pump-down)
Input Pin config
Start Logging System
- init errors
Start Logging System
- init errors
Perform selfChecks()
- read switch states
- verify valve positioning
Read switch states
Yes
Failed check?
Log serial
No
updateLCD()
Verify 972b Status
- Transducer status check (@xxxT?;FF)
LCD StartupMsg
Retry selfChecks()
Pressure Sensor Configuration
- Config safety relay setpoint
- Enable safety relay
- Verify output type, precision
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Raise ERROR + updateLCD()
“R”
Raise MICROPIRANI_FAILURE
“O”
Pressure dose setpoint exceeded
Sensor OK
SP1 relay error should be persistent and not clearable.
Loop back up to top read Switch States
Raise PRESSURE_DOSE_WARNING
Clear any 972b status errors from LCD queue
Set units to mbar
Delay
Add checks to parse for other errors besides NACK. i.e. incomplete msgs.
Request
Initial Pressure
ACK?
No
No
Retry limit
exceeded?
Delay
Yes
Yes
ACK?
No
Yes
No
Retry limit
exceeded?
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Yes
Yes
Yes
Raise ERROR
Pressure NACK
Clear any pressure NACK error from
LCD queue
Set user tag
Request Initial Pressure reading
Display Pressure
TODO: update error count based on response. Clear prior errors before moving forward
Check if initial pressure approx.
~1013 mbar
Verify Pressure Safety Relay OPEN?
Yes
Pressure below
1E-6 mbar?
(HI_VAC)
Raise UNEXPECTED_PRESSURE_ERROR
No
Within
Threshold?
Power cut to
High voltage
electronics
No
(PUMP_DOWN)
Raise PRESSURE_RELAY_ERROR Expected: OPEN
Actual: CLOSED
yes
No
OPEN?
Verify setpoint relay direction: BELOW
If previously had NACK error, and now ACKs, clear previous error
Yes
Delay
Clear any PRESSURE_RELAY_ERROR from LCD queue
Verify Pressure Safety Relay CLOSED
ACK?
No
No
Retry limit
exceeded?
Yes
Yes
Raise SAFETY_RELAY_ERROR Expected: CLOSED
Actual: OPEN
Yes
Raise PRESSURE_NACK_ERROR
No
CLOSED?
Asynchronous LCD refreshes every 0.5 second with latest pressure and error info
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
No
Yes
Yes
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Set Safety relay setpoint: 1.00E-6 mbar
Yes
No
Delay
ACK?
No
No
Retry limit
exceeded?
Clear any TURBO_ROTOR_ON_WARNING
Raise VALVE_CONTENTION_ERROR
Yes
Contention of valve
positions?
TURBO_ROTOR_ON?
Yes
Yes
(reading open and
closed simultaneously)
Yes
Raise PRESSURE_NACK_ERROR
No
No
Clear any PRESSURE_NACK_ERROR
from LCD queue
Clear any VALVE_CONTENTION_ERROR from LCD queue
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
Yes
Vent Valve Switch
OPEN?
Query Pressure safety relay status
Raise
ARGON_GATE_VALVE_ERROR
No
Argon Gate valve
OPEN?
(closed)
Raise SAFETY_RELAY_ERROR
No
ENABLED?
Yes
Yes
Clear any ARGON_GATE_VALVE_ERROR from LCD queue
- Sequence of venting actions to ATM
1. turbo rotor off
2. PUMPS_POWER_OFF
3. Vent valve switch open (wait for turbomolecular pump to decelerate)
4. Gate valves closed (argon and turbo)
5. Once, completed, and pressure within limit,
Raise
TURBO_GATE_VALVE_WARNING
No
Turbo Gate valve
OPEN?
(closed)
Yes
Clear any TURBO_GATE_VALVE_WARNING from LCD queue
Raise
VENT_VALVE_OPEN_ERROR
Yes
Vent valve
switch OPEN?
No
Clear any VENT_VALVE_OPEN_ERROR from LCD queue
Serial log
ValveConfig check OK
Vacuum loss error handling threshold
Raise status
UNSAFE_FOR_HV
Yes
Pressure above
1E-4 ?
“unsafe for HV” warning cycles with error messaging
No
No
No
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
Yes
Yes
Clear any
TURBO_ROTOR_ON_WARNING
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
Delay
Remove this
(only check once, conditional on the current pressure)
TODO:
Swap vertical position of turbo rotor check with “pressure above 1E-4”

---
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Yes
No
Log serial + updateLCD()
Functional Tests
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Command Relay OFF
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
no
Transition to: DEBUG_STATE
No
ACK?
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Retry limit
exceeded?
Timeout?
Timeout?
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Yes
Yes
yes
Yes
No
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Transition to: DEBUG_STATE
“M”
Cold cathode failure
MicroPirani failure
“C”
Disable setpoint relay
status?
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Within
Threshold?
~1013 mbar
Log error + updateLCD()
No
“R”
Log error + updateLCD()
“O”
yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
Sensor OK
Yes
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
No
No
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Yes
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
Yes
ACK?
No
no
Log serial
Transition to: IDLE_STATE
Retry limit
exceeded?
IS THIS 10%
No
Setpoint
below 1 atm?
No
Yes
Log + updateLCD()
Retry limit
exceeded?
yes
Yes
updateLCD()
Yes
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Retry limit
exceeded?
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
Within
Threshold?
~1 atm
ABOVE
Yes
Direction?
Greater than
 +/-10% of setpoint?
Log serial
yes
No
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
updateLCD()
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay