Power On
setup()
Initialize Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Input Pin initialization
Read switch states
Pressure Sensor Configuration
- Includng safety relay setpoint
- Enable safety relay
- Verify output type, precision
Set units to mbar
Delay
ACK?
No
Retry limit
exceeded?
Yes
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Set user tag
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Set Safety relay setpoint: 1.00E-4 mbar
BW says change this to
1.00E-4 mbar
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Verify Pressure safety relay enabled
    ACK?
Raise PRESSURE_NACK_ERROR
Request Initial Pressure reading
    ACK?
Check if initial pressure approx.
~1013 mbar
Loop NACK retries back up to re-read
Pressure, not just to previous command
Raise PRESSURE_NACK_ERROR
03/28:
Deactivate this after
10 minutes
Raise UNEXPECTED_PRESSURE_ERROR
Within
Threshold?
yes
Verify setpoint relay direction: BELOW
Raise PRESSURE_NACK_ERROR
Clear any PRESSURE_NACK_ERROR
from LCD queue
Verify
errorcount == 0
972b
Error?
Log status
SETUP COMPLETE
Transition to
NORMAL OPERATION Loop
VTRX-200 Software Flow – Setup
03-2024 (standard/argon pump-down)

---
NORMAL_OPERATION
(standard/argon pump-down)
[enter]
Read system status
Verify 972b Status
- Transducer status check (@xxxT?;FF)
“M”
Cold cathode failure
MicroPirani failure
“C”
status?
Raise ERROR + updateLCD()
“R”
Raise MICROPIRANI_FAILURE
“O”
Pressure dose setpoint exceeded
Sensor OK
Raise PRESSURE_DOSE_WARNING
Clear MicroPirani or ColdCathode status errors from LCD queue
Request
Initial Pressure
RETRY LIMIT:3
Raise ERROR
Pressure NACK
Clear any active PRESSURE_NACK_ERROR from LCD error queue
If previously sensor NACK’s, and now ACKs, clear previous error
Display Pressure
Power cut to
High voltage
electronics
Update pressure in this screenshot
Verify Pressure Safety Relay OPEN?
Pressure below
1E-4 mbar +
threshold?
(HI_VAC)
“unsafe for HV” warning cycles with error messaging if present
If RELAY_ERROR has been
exists for >1 sec,
Make the error persistent
(PUMP_DOWN)
Raise PRESSURE_RELAY_ERROR Expected: OPEN
Actual: CLOSED
OPEN?
Raise status
UNSAFE_FOR_HV
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Verify Pressure Safety Relay CLOSED
Clear any UNSAFE_FOR_HV status
Raise SAFETY_RELAY_ERROR Expected: CLOSED
Actual: OPEN
Raise VALVE_CONTENTION_ERROR
CLOSED?
Contention of valve
positions?
(reading open and
closed simultaneously)
If this error has been existing for >1 sec,
Make the error persistent
Clear any active PRESSURE_RELAY_ERROR from LCD queue
Clear any active VALVE_CONTENTION_ERROR from LCD queue
Raise VALVE_CONTENTION_ERROR
Contention of valve
positions?
(reading open and
closed simultaneously)
Clear any VALVE_CONTENTION_ERROR from LCD queue
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
PUMPS_POWER_ON?
TURBO_ROTOR_ON?
TURBO_ROTOR_ON?
PUMPS_POWER_ON?
Clear any TURBO_ROTOR_ON_WARNING
Clear any TURBO_ROTOR_ON_WARNING
Clear any
TURBO_ROTOR_ON_WARNING
Clear any
TURBO_ROTOR_ON_WARNING
Raise
TURBO_ROTOR_ON_WARNING
Expected: Turbo OFF
Actual: Turbo ON
TBD Delay ms
TBD Delay ms
Notes
- This system is wholly dependent on operator actions.
- No automated hardware interlocks prevent operator from progressing to next step if the current step isn’t completed properly.
Primary Directives
(agnostic to what operator is doing)
- display pressure
- display switch malfunction (contention)
- HV relay state verification
Displays warnings for the following conditions:
- TURBO_ROTOR_ON while PUMPS_POWER OFF
- UNSAFE_FOR_HV messages (pressure above high-voltage safety threshold)
- Sequence of venting actions to ATM
  (for operator)
1. turbo rotor off
2. PUMPS_POWER_OFF
3. Vent valve switch open (wait for turbomolecular pump to decelerate)
4. Gate valves closed (argon and turbo)
5. Once completed, verify pressure within 1 atm threshold
VTRX-200 Software Flow – Main loop

---
VTRX-BTEST-040
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Set user tag
Watchdog timer
Watchdog timer
Verify Pressure sensor Relay 1 is disabled (EN1?) (default)
Timeout?
Log serial + updateLCD()
Functional Tests
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Log error (serial + LCD)
As expected?
(de-energized)
Command Relay OFF
Log serial + updateLCD()
TODO: VTRX-BTEST-050
Transition to: DEBUG_STATE
VTRX-BTEST-020
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Verify 972b comms
- Set user tag
Test serial reading from pressure gauge using arduino at 1 atm and roughing
Log error + updateLCD()
Request Pressure
no
Transition to: DEBUG_STATE
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
no
Timeout?
Timeout?
Log serial + updateLCD()
Verify pressure within threshold for 1 atm
Log error + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Transition to: DEBUG_STATE
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Reflect reading not as expected
Log error + updateLCD()
Log error + updateLCD()
Log error (serial + LCD)
As expected?
(de-energized)
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Query safety relay enable value (SP1?)
Command Relay OFF
Log serial + updateLCD()
Setpoint
Below 1 atm?
Verify pressure units: mbar
Transition to: IDLE_STATE
Log + updateLCD()
Loop
Timeout?
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Query setpoint hysteresis value
Log error + updateLCD()
Request Pressure
no
Query setpoint switch value (SP1?)
Transition to: IDLE_STATE
no
Log serial
Transition to: IDLE_STATE
IS THIS 10%
Setpoint
below 1 atm?
Log + updateLCD()
updateLCD()
Verify reading within threshold for 1 atm
Log error + updateLCD()
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Set SH1 hysteresis: 10%
Query setpoint relay direction
Transition to: IDLE_STATE
Query setpoint hysteresis value
Transition to: IDLE_STATE
Perform calibration
no
Command setpoint relay direction: BELOW
~1 atm
ABOVE
Direction?
Greater than
 +/-10% of setpoint?
Log serial
BELOW
updateLCD()
Print pressure
Enable relay
Verify safety delay enabled**
Set SH1 hysteresis
Query relay state to verify
Log SUCCESS +
updateLCD()
Enable 972b setpoint
Log serial + UpdateLCD()
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Report triggered
Log serial + updateLCD
Transition to IDLE_STATE
Verify Relay Enabled
TODO: above/below
Then pump down to see if we can -detrigger the relay