Power On
setup()
init SystemStates
set currentState = INITIALIZATION_STATE
INITIALIZATION STATE
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
Main Loop
Configure input pins
switch (currentState)
Start Logging System
- init errors
Start Logging System
- init errors
Perform selfChecks()
- read power states
- verify valve positions
- set LED indicators
- check pressure is 1 atm
De-Energized State
Yes
Failed check?
Log serial
Actions to return to
de-energized config
- Pressure Gauge OFF
- Pumps power OFF
- close gas valves
No
updateLCD()
updateLCD()
Retry selfChecks()
updateLabVIEW()
updateLCD()
Safe_State
Error State
Log Error
Triggered by Manual Shutdown or Error?
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
VTRX-BTEST-040
Assume most functions log at INFO level
Explici
STANDARD_PUMP_DOWN STATE
Verify 972b comms
- Set user tag and display response
Verify 972b pressure
- confirm 1 atm
No
Watchdog timer
Watchdog timer
Verify Relay 1 is disabled (EN1?)
Timeout?
Timeout?
Yes
Yes
Check Valve Pos.
- Gate valve switches: OPEN
- Vent valve switch: CLOSED
No
Log serial + updateLCD()
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Yes
Disable setpoint relay
Perform selfChecks()
- read power states
- verify valve positions
- set LED indicators
Query Pressure Relay 1 (SS1?) status
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
Failed check?
Log serial
Yes
No
updateLCD()
updateLCD()
Log serial + updateLCD()
updateLCD()
Transition to: IDLE_STATE
Retry selfChecks()
updateLabVIEW()
Log serial
Query setpoint switch value (SP1?)
updateLCD()
No
Setpoint
below 1 atm?
Set SP1 setpoint below ambient
e.g. 5.00E+1 torr
No
Timeout?
Yes
Yes
Query setpoint hysteresis value
Watchdog timer
Watchdog timer
Yes
Greater than
 +/-10% of setpoint?
Log serial
No
Transition to: IDLE_STATE
Transition to: IDLE_STATE
updateLCD()
Enable setpoint
Set SH1 hysteresis