Power On
setup()
INITIALIZATION STATE
TODO:
STANDARD_PUMP_DOWN STATE
IDLE_STATE
init SystemStates
Establish Comms
- Serial: program + log
- Serial1: LabVIEW
- Serial2: Pressure Gauge
set currentState = INITIALIZATION_STATE
Log + updateLCD()
Verify 972b pressure
- confirm 1 atm
Input Pin config
Check Valve Pos.
- Gate valve switches: OPEN
- Vent valve switch: CLOSED
Start Logging System
- init errors
Start Logging System
- init errors
Check Valve Pos.
- Gate valve switches: OPEN
- Vent valve switch: CLOSED
Main Loop
Perform selfChecks()
- read power states
- verify valve positioning
- set LED indicators
switch (currentState)
Perform selfChecks()
- read power states
- verify valve positions
- set LED indicators
Perform selfChecks()
- read power states
- verify valve positions
- set LED indicators
Yes
Failed check?
Log serial
Yes
Failed check?
Log serial
No
Yes
Failed check?
Log serial
updateLCD()
No
StartupMsg + updateLCD()
updateLCD()
No
updateLCD()
Log serial + updateLCD()
Retry selfChecks()
updateLCD()
updateLabVIEW()
Retry selfChecks()
Transition to REMOTE_CONTROL
State
updateLabVIEW()
Request for
REMOTE ctl?
Transition to
TEST_STATE/PUMP_DOWN
No
Read Pressure
Log pressure + updateLCD()
Delay
TODO: VTRX-BTEST-050
VTRX-BTEST-020
VTRX-BTEST-040
Test serial reading from pressure monitor in Labview sub-VI from pressure gauge at 1 atm
Test serial reading from pressure gauge using arduino at 1 atm
Test pressure safety relays reading from pressure gauge using arduino at 1 atm
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
Verify 972b Status
- Transducer status check (@xxxT?;FF)
Verify 972b comms
- Transducer status check (@xxxT?;FF)
- Set user tag
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
“M”
Cold cathode failure
MicroPirani failure
“C”
Timeout?
Timeout?
status?
Yes
No
Watchdog timer
Watchdog timer
Query Pressure sensor if Relay 1 is disabled (EN1?) (default)
Log error + updateLCD()
Timeout?
Timeout?
“R”
No
Log serial + updateLCD()
Log error + updateLCD()
Yes
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
“O”
Yes
Transition to: IDLE_STATE
Pressure dose setpoint exceeded
No
Log serial + updateLCD()
Sensor OK
Off?
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Disable setpoint relay
Transition to: IDLE_STATE
Query Pressure Relay 1 (SS1?) status
Yes
Disable setpoint relay
Query Pressure Relay 1 (SS1?) status
Set pressure units: mbar
No
Log error (serial + LCD)
As expected?
(de-energized)
Yes
No
Log error (serial + LCD)
As expected?
(de-energized)
Command Relay OFF
Log error + updateLCD()
Request Pressure
Log serial + updateLCD()
Yes
no
Command Relay OFF
Transition to: IDLE_STATE
Transition to: IDLE_STATE
Log serial + updateLCD()
ACK?
no
Retry limit
exceeded?
Transition to: IDLE_STATE
Yes
yes
Yes
Verify reading within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Query setpoint switch value (SP1?)
No
Transition to: IDLE_STATE
Log error + updateLCD()
Request Pressure
No
Setpoint
below 1 atm?
Perform calibration
no
Log + updateLCD()
no
Retry limit
exceeded?
Within
Threshold?
~1 atm
Transition to: IDLE_STATE
Yes
ACK?
Yes
no
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
yes
Retry limit
exceeded?
Transition to: IDLE_STATE
Query setpoint hysteresis value
Yes
Print pressure
yes
Yes
Verify pressure within threshold for 1 atm
Log error + updateLCD()
No
Retry limit
exceeded?
Yes
Greater than
 +/-10% of setpoint?
Log SUCCESS +
updateLCD()
Log serial
Transition to: IDLE_STATE
No
updateLCD()
Perform calibration
Within
Threshold?
~1013 mbar
Configure setpoint safety delay**
No
Transition to: IDLE_STATE
Set SH1 hysteresis
yes
Query setpoint switch value (SP1?)
Enable 972b setpoint
No
No
Setpoint
Below 1 atm?
Log + updateLCD()
Log serial + updateLCD
Loop
Timeout?
Yes
**The setpoint safety delay function requires 5 continuously measurements that exceeds setpoint value before the relay is tripped. This feature prevents that noise or pressure pulse can trig the relay.
Yes
Set SP1 setpoint below ambient
e.g. 10^-1 millibar
Verify Relay Enabled
Watchdog timer
Query setpoint hysteresis value
Watchdog timer
Yes
Greater than
 +/-10% of setpoint?
Log serial
Transition to: IDLE_STATE
No
TODO: above/below
updateLCD()
Set SH1 hysteresis: 10%
Query setpoint relay direction
Command setpoint relay direction: ABOVE
BELOW
Direction?
ABOVE
Enable relay
Query relay state to verify
Log serial + UpdateLCD()
TODO: VTRX-BTEST-900